## Tabelle1
The accuracy (is the total ozone value calculated the real one) is one part, the precision is how well can the measurement be repeated is the other. Some of errors are in the instrument and its operation and can be both, others are in the algorithm and effect the accuracy.
item | error source | in detail | affected instrumental and other quantities | amount | random/systematical (Precision vs Accuracy?) | origin | comments | Bob Evans Comments:
1 | wavelength adjustment | incorrect Q-table, uncertainty in instrumental temperature, especially during rapid T-changes | absorption coefficients (effective) | 1% | random | Basher 1982 | consistent with PTB results | If the instrument is operated consistantly on the wrong wavelengths, then this is a accuracy issue. Calibration against a standard or by langley plot will not fix this completely
2 | bandwidth effect | optical misalignment | absorption coefficients (effective) | - | random | Basher 1982 | included in item 1
3 | ozone absorption | inaccurate determination of ozone cross sections in the lab | absorption coefficients (nominal and effective) | 1% | random | Basher 1982 (prospects for future after 1982) | currently relevant: difference between BP and DMB coefficients | (Accuracy) For Dobson and Brewer, The ozone cross-section database determined by Institute of Environmental Physics, University of Bremen seems best.
4 | temperature ozone layer | assumed to be -46o C, mostly lower temperatures (Ozone weighted temperature -- ozone profile convoluted with temperature profile) | absorption coefficients (effective) | up to -1% | mostly systematic |  | seasonally depending: winter maximum | This error (accuracy) is latitude dependent. Southern Hemisphere errors can be quite large during depletion events. Moon observations can incorrect, as the temperture of the statosphere is very cold.
5 | atmospheric scattering | determination of Rayleigh coefficients | Rayleigh coefficients | < 0.1% | random | own estimation | valid for routine observations in the double wavelengths pair modes AD und CD
6 | particle scattering | particel scattering in hazy atmosphere | aerosol scattering coefficients | < +0.2% | systematic | Basher 1982 | only relevant for single wavelength pair observation (+3%) at high turbidity
7 | direct intercomparison | overall uncertainty of the observations (depends on weather situation and skill of operator) | ETCs | 1% | random | Basher 1982 / own experience | seems reasonable from my experience during DICs at the RDCC-E
8 | Langley method for abs. cal. of reference |  | ETCs | 5% (0.5%) | random | Basher 1982 / Bob | not consistent with Bob Evans' findings (his Langley results with D083 mostly within 0.5%) | The +/-0.5 is the repeatibility. The accuracy of the determined ETC is not well known, but as our results when compared to Brewer Langley plots at Izana show good agreement.
9 | spectral stability | changes of the spectral sensitivy: deterioration of cobalt filter, PMT-changeing (ageing, changing position, non-linearity) etc. | ETCs | 0.5% | random | Basher 1982 / own experience | Basher claims errors up to 1.5% or even higher. This is only justified, when SL test results are not applied in correcting the ETCs
10 | relative optical path through the ozone layer | calculation of mue uses fix height of ozone layer, depending on latitude; climatological mean, can vary with weather situation | Mue | 0 to 1% | random | Basher 1982 / own estimations | depending on solar zenith angle (1% at mue = 4 and delta in height 4km) | Langley Plot is dependent on this factor.
11 | direction of measurement | Total ozone column means ozone between instrument and zenith, measurements, however are done to the sun | TOC as external value | 1 - 2 % | random | own estimation | depending on solar zenith angle and gradient in stratospheric ozone | 2-5% variablity ZS versus ADDS -- there might be a offset for cloudy zenith due to the conditions under which the zenith algorthim is determined (days with both zenith and DS observations.)
12 | wedge calibration | difficulties during wedge calibration to determine R-G-tables as base for the needed R-N-tables | N-values from raw data R-values | 0.5% | random | Basher 1982 | some optical wedges have steep gradient, difficult to measure
13 | stray light (internal) | misdirected light of unwanted longer wavelengths in a Dobson | R-, N-values | < -0.1% | systematic | Basher 1982 / Bob | relevant only at low sun (mue > 3 - 3.5) or high ozone value, when signal in measured wavelengths very small | The effect of Internally scattered light is dependent on the input intensity -- which is dependent on TOC*Mu (Slant Path ozone). Umkehr measurements are also affected by this "out of band" light. So far, the magnitude of the internally scattered light has not been accurately measured.
14 | stray light (external) = skylight | stray light around the sun disc with different spectral composition; Dobson has a rather large field of view with 8 degree | R-, N-values | < -1% at mue < 3 | systematic | Basher 1982 / own experience | depends on turbidity and limits max mue-values normally less than 3.2 (up to 4 at very clear sky conditions and high altitudes | This is difficult evaluate, as there are many factors as Ulf noted, and overlaps in the region of slant path ozone where internally scattered light becomes a problem. Focused image observations work better, but require skilled operators.
15 | solar constancy | affects ETCs, when N0-values changed | R-N-table | <0.1% | random | Basher 1982
16 | interference with other atmospheric constituents | SO2 (either from power plants or volcanic eruptions) can influence TOC observations | TOC as external value | normally < +0.1%; up to +5% possible | systematic | Basher / own experience | large interference uncommon, since power plants are desulfurized, last incident happened during eruption of Eyjafjallajökull (about 5 D.U. = +1.5%)
17 | operational precision | depends on skill/experience of operator | TOC as external value | < 1% | random | Basher / own experience | caused e.g. by misalignment of Dobson to the sun, wrong high voltage setting etc.; can increase up to 2% | This is a very important part of the system, and often overlooked by users of the data.

## Tabelle3
u_rel_wv
u_rel_xs
u_rel_temp
u_rel_btest
u_rel_beta1
u_rel_beta2
u_rel_P
u_rel_delta
u_rel_sec
u_rel_etc
u_rel_mu
u_rel_m
u_rel_wedge
u_rel_straylight
u_rel_wedge
u_rel_interference
u_rel_operat_impre